## Summary
ITEM | FUME CABINETS, UNIVERSITY OF VENDA - MECHANICAL WORKS | AMOUNT
1 | P & Gs @ 8%
2 | FUME CABINETS
3 | SUB TOTAL BEFORE CONTINGENCIES
4 | CONTINGENCIES @ 5%
5 | SUB TOTAL AFTER CONTINGENCIES
6 | VAT
7 | GRAND TOTAL

## Preliminary and General
FUME CABINETS, UNIVERSITY OF VENDA - MECHANICAL WORKS
ITEM | DESCRIPTION | UNIT | QTY | RATE | AMOUNT
PRELIMINARY & GENERAL
Site Facilities
1 | Site camp, accommodation and office | Sum | 1 |  |  | 5
2 | Security | Sum | 1 |  |  | 5
Insurances
3 | Works Insurance | Sum | 1 |  |  | 5
4 | Additional insurance for long lead-time items | Sum | 1 |  |  | 3
5 | Storage for long lead-time items | Month | 1 |  |  | 3
6 | Workshop drawings | Sum | 1 |  |  | 5
7 | Builders work drawings | Sum | 1 |  |  | 2.5
8 | Allowance for drawings in electronic format | Sum | 1 |  |  | 1.5
9 | Operation and maintenance manuals | Sets | 4 |  |  | 5
10 | One year written guarantee and maintenance for 12 months | Sum | 1 |  |  | 10
Health and Safety
11 | Safety, Health and Environmental Management | Sum | 1 |  |  | 15
12 | Personal Protective Clothing | Sum | 1 |  |  | 10
13 | Training of client's staff | Sum | 1 |  |  | 5
14 | Contract management | Sum | 1 |  |  | 20
15 | Painting and marking | Sum | 1 |  |  | 2.5
16 | Clearing of site facilities upon completion | Sum | 1 |  |  | 2.5
90
TOTAL CARRIED TO SUMMARY | [5 empty] | 0

## Life Sciences
FUME CABINETS, LIFE SCIENCES DEPARTMENT, UNIVERSITY OF VENDA - MECHANICAL WORKS
ITEM | DESCRIPTION | UNIT | QTY | RATE | AMOUNT
1 | LIFE SCIENCE LAB FUME CABINETS
1.1 | Supply, delivery, removal and Installation of fume cabinets complete with ancilliaries
1.1.1 | 1.3m free flow fume cupboard full polypropylene construction. Unit complete with fully bunded base constructed from laboratory grade phenolic resin G-Lab, black in colour. Unit with integrated lighting, vertical sliding polycarbonate 6mm sash on counter weight. Baffle to be removable for easy cleaning. External size 1300 x 750 x 1520mm excluding stand, overall height with stand@ 2420mm Internal working space 1000 x 630mm. Unit will have removable side service panels as well as top panel. | No. | 13 |  |  |  |  | 47
1.1.2 | Fume cupboard light switch | No. | 13 | Incl. | 0
1.1.3 | 15A plugs | No. | 26 | Incl. | 0
1.1.4 | Broen water front control valve to operate with water outlet nozzle inside the unit. Complete with polypropylene drip cup and fittings | No. | 13 | Incl. | 0
1.1.5 | Fume cupboard stand of 32 x 32mm epoxy coated tubular frames complete with connecting tie bars coated to client’s choice | No. | 13 | Incl. | 0
1.1.6 | Set under bench cabinet under the unit consisting of one 900mm wide and one 500mm wide units | No. | 13 | Incl. | 0
1.1.7 | Polypropylene fume extraction ducting manifolds to connect two 1.3m fume cupboards to chemical extractor fan on the roof including all bends, segments, bracketing and flashing |  |  | Sum
1.1.8 | 1.5m free flow fume cupboard full polypropylene construction. Unit complete with fully bunded base constructed from laboratory grade phenolic resin G-Lab, black in colour. Unit with integrated lighting, vertical sliding polycarbonate 6mm sash on counter weight. Baffle to be removable for easy cleaning. External size 1500 x 750 x 1520mm excluding stand, overall height with stand@ 2420mm Internal working space 1200 x 630mm. Unit will have removable side service panels as well as top panel. | No. | 44 |  |  |  |  | 237000
1.1.9 | Complete With:
1.1.10 | Fume cupboard light switch | No. | 44 | Incl. | 0
1.1.11 | 15A plugs | No. | 88 | Incl. | 0
1.1.12 | Broen water front control valve to operate with water outlet nozzle inside the unit. Complete with polypropylene drip cup and fittings | No. | 44 | Incl. | 0
1.1.13 | Fume cupboard stand of 32 x 32mm epoxy coated tubular frames complete with connecting tie bars coated to client’s choice | No. | 44 | Incl. | 0
1.1.14 | Set under bench cabinet under the unit consisting of one 900mm wide and one 500mm wide units | No. | 44 | Incl. | 0
1.1.15 | Polypropylene fume extraction ducting manifolds to connect 1.5m fume cupboards to chemical extractor fan on the roof including all bends, segments, bracketing and flashing |  |  | Sum | 105000
1.1.16 | Polypropylene fume extraction ducting manifolds to connect 1.3m fume cupboards to chemical extractor fan on the roof including all bends, segments, bracketing and flashing |  |  | Sum | 225000
1.1.17 | Perchloric acid fume hoods 1,5m free flow fume cupboard full polypropylene construction Unit complete with a fully bunded base constructed from laboratory grade phenolic resin G-Lab, black in color. Unit with integrated lighting, vertical sliding polycarbonate 6mm sash on the counterweight. Baffle to be removable for easy cleaning. External size 1500 x 750 x 1520mm excluding stand, overall height with stand @ 2420mm Internal working space 1200 x 630mm. Unit will have removable side service panels as well as top panel. Each unit is complete with the following: Fume cupboards light switch 15A plugs Set heat sheilds for use with hot plates Fume cupboard stand for 1,5m fume cupboard Set fume cupboards wash down sprays, front controlled to operate the spray bar inside the fume cupboard Broen Lab water front control valve to operate the outlet nozzel inside the fume cupboard complete with polypropylene drip cup and anti-siphon bottle trap Broen Lab Compressed Air front control valve to operate the outlet nozzle inside the fume cupboard complete with all fittings Broen Lab LP Gas front control valve to operate the outlet nozzle inside the fume cupboard complete with all fittings Broen Lab Vacuum front control valve to operate the outlet nozzle inside the fume cupboard complete with all fittings | No. | 2
1.1.18 | Fume Extraction System 1 x Complete set polypropylene extraction ducting connecting two 1500mm fume cupboards in series to external centrifugal chemical extractor fan on the exterior and exhaust up and away from the building approximately 5-6m 1 x Fan chassis 1 x CF 300/14/380V PVC centrifugal extractor fan complete with ducting all in rigid Polyprop | Sum |  | 54500 | 54500
1.1.19 | CF 300/14/380V PVC centrifugal extractor fan complete with ducting all in rigid Polyprop with 1.5kw 380v motor with weather proof isolator | No. | 20
1.1.20 | CF 300/14/380V PVC centrifugal extractor fan complete with ducting all in rigid Polyprop with 1.1kw 380v motor with weather proof isolator | No. | 2
1.1.21 | CF 250/14/380V PVC centrifugal extractor fan complete with ducting all in rigid Polyprop with 0.55kw 380v motor with weather proof isolator | No. | 5 |  |  | 29 |  | 384500
1.1.22
1.1.23 | Supply, deliver and install
1.1.24 | Oil free Vacuum Pump, one duty, one standby, Flow Rate: 65m3/hr, Power Consumption: 1.1kW, 1PH, 240V, 50Hz | No. | 2
1.1.25 | Vacuum Receiver Tank, Capacity: 500L | No. | 1 | Incl.
1.1.26 | Vacuum Control Panel | Sum | 1 | Incl.
1.1.27 | Pressure Gauge | No. | 1 | Incl.
1.1.28 | Vacuum Trap | No. | 1 | Incl.
1.1.29 | Isolation Valves | No. | 4 | Incl.
1.1.30 | Oil Free Piston type compressor, one duty, one standby Flowrate: 20m3/hr Power Consumption: 1.5kW, 1PH, 240V, 50Hz | No. | 2
1.1.31 | Refrigerant Type Compressor Air Drier, one duty, one standby, Refrigerant Type: R410A, Flowrate: 2.25L/s ,Pressure: 700kPa, Power Consumption: 1.62KW, 1PH, 240V, 50Hz | No. | 2
1.1.32 | Compressor Air Receiver Tank, Capacity: 1000L | No. | 1
1.1.33 | Oil Trap Filter | No. | 1
1.1.34 | Water Trap Filter | No. | 1
1.1.35 | Isolation Valves | No.
1.1.36
1.1.42 | Supply, delivery and installation of high purity equipement, complete with all interconnecting 316L Stainless Steel Piping and fittings
1.1.43 | High purity installation for N2/Helium supply | No. | 2
1.1.44
1.1.45 | High purity installation for H2/N2/Synthetic Air/Argon/Helium/P10 supply | No. | 8
1.1.46
1.1.47 | LP Gas, Vacuum and compressed air pipes as per SABS 460 class 2 pipe line
1.1.48
1.1.49 | 35mm copper pipe for reticulation network | m | 434
1.1.50
1.1.51 | 35mm 90 deg elbow | No. | 152
1.1.52
1.1.53 | 35mm x 35mm x 22mm Tee | No. | 131
1.1.54
1.1.55 | 35mm x 35mm x 15mm Tee | No. | 46
1.1.56
1.1.57 | 35mm x 35mm x 10mm Tee | No. | 40
1.1.58
1.1.59 | 22mm copper pipe for reticulation network | m | 168
1.1.60
1.1.61 | 22mm 90 deg elbow | No. | 59
1.1.62
1.1.63 | 22mm x 22mm x 15mm Tee | No. | 51
1.1.64
1.1.65 | 22mm x 22mm x 10mm Tee | No. | 18
1.1.66
1.1.67 | 15mm copper pipe for reticulation network | m | 224
1.1.68
1.1.69 | 15mm 90 deg elbow | No. | 78
1.1.70
1.1.71 | 15mm x 15mm x 15mm Tee | No. | 68
1.1.72
1.1.73 | 15mm x 15mm x 10mm Tee | No. | 24
1.1.74
1.1.75 | 10mm copper pipe for reticulation network | m | 413
1.1.76
1.1.77 | 10mm 90 deg elbow | No. | 144
1.1.78
1.1.79 | 10mm x 10mm x 10mm Tee | No. | 87
1.1.80
1.1.81 | 35mm lockable isolating valves | No. | 7
1.1.82
1.1.83 | 22mm Lockable Isolating Valves | No. | 5
TOTAL
TOTAL
1.1.89 | [6 empty] | 16
1.1.90 | 48 kg LP Gas cylinders | No | 48
1.1.91 | LP Gas Cage | No | 3
1.1.92 | 25mm LP Gas Regulators | No | 5
1.1.93 | 10mm aluminium flexible hoses | No | 56
1.1.94
1.1.95 | Water pipes as per SABS 460 class 1 pipe line
1.1.96
1.1.97 | 35mm copper pipe for reticulation network | m | 217
1.1.98
1.1.99 | 35mm 90 deg elbow | No. | 76.3
1.1.100
1.1.101 | 35mm x 35mm x 22mm Tee | No. | 65
1.1.102
1.1.103 | 35mm x 35mm x 15mm Tee | No. | 23
1.1.104
1.1.105 | 35mm x 35mm x 10mm Tee | No. | 20
1.1.106 | 22mm copper pipe for reticulation network | m | 84
1.1.107
1.1.108 | 22mm 90 deg elbow | No. | 30
1.1.109
1.1.110 | 22mm x 22mm x 15mm Tee | No. | 26
1.1.111
1.1.112 | 22mm x 22mm x 10mm Tee | No. | 9
1.1.113
1.1.114 | 15mm copper pipe for reticulation network | m | 112
1.1.115
1.1.116 | 15mm 90 deg elbow | No. | 40
1.1.117
1.1.118 | 15mm x 15mm x 15mm Tee | No. | 34
1.1.119
1.1.120 | 15mm x 15mm x 10mm Tee | No. | 12
1.1.121
1.1.122 | 10mm copper pipe for reticulation network | m | 207
1.1.123
1.1.124 | 10mm 90 deg elbow | No. | 73
1.1.125
1.1.126 | 10mm x 10mm x 10mm Tee | No. | 62
1.1.127
1.1.128 | 35mm lockable isolating valves | No. | 4
1.1.129
1.1.130 | 22mm Lockable Isolating Valves | No. | 2
1.1.131
1.1.132 | 15mm Lockable Isolating Valves | No. | 3
1.1.133
1.1.134 | 10mm Lockable Isolating Valves | No. | 16
1.1.135
1.1.136
1.1.137 | Drainage pipes as per Class 6 PVC pipe line
1.1.138
1.1.139 | Supply, deliver, off-load, store on site and install of Class 6 PVC piping including all fittings and other accessories for the drainage plumbing work , as required
1.1.140
1.1.141 | 110mm diameter | No | 175
1.1.142
1.1.143 | 50mmdiameter | No | 193
1.1.144
1.1.145 | 35 diameter | No | 46
1.1.146
1.1.147 | 28 diameter | No | 42
1.1.148
1.1.149 | 22 diameter | No | 123
1.1.150
1.1.151 | Supply, delivery, installation and commissioning of 50mm bend wiith Cleaning eye heel or (Rodding eye) | No | 14
1.1.152 | Supply, delivery, installation and commissioning of gully trap complete excluding building work (Gully Base by Main contractor) | No | 17
1.1.153 | Supply, delivery, installation and commissioning of 110mm Vent Valve | No | 20
1.1.154
1.1.155 | Laminar flow 1 x IHAIHR1200-IIA2-NRCS, BIOSAFETY CABINET CLASS II TYPE A2 4FT INCL UV LIGHT AND BASE STAND Working Voltage&Frequency(V/Hz) 220V 50Hz, Power (VA) 1500, Interior Dimensions(W*D*H) 1220*620*650 mm, Exterior Dimensions(W*D*H) 1380*845*2160mm | No | 1
1.1.156 | BENCHES AND GRANITE TABLES
1.1.157 | Additional laboratory structures such as granite tables in 6x laboratories b. Additional shelving space in 6 x laboratories | Sum |  | 125500 | 125500 |  |  |  | .
1.1.158 | 180 degree double LPG bench taps. 180 degree double vacuum bench taps. All upright support from ceiling height to island benches.All interconnecting copper piping. e. All testing and commissioning. | Sum |  | 275000 | 275000
1.1.159 | Instruments Complete high purity installation for H2/N2/Synthetic Air/Argon/Helium/P10 supply to 4x GC, 1x TGA, 1x ICP, 1x AA, 1x XRF. Including testing and commissioning | No | 1
Carried to summary

## Environmental Sciences
FUME CABINETS, ENVIRONMENTAL SCIENCES DEPARTMENT, UNIVERSITY OF VENDA - MECHANICAL WORKS
ITEM | DESCRIPTION | UNIT | QTY | RATE | AMOUNT
2 | ENVIRONMENTAL SCIENCES FUME CABINETS
2.1 | Supply, delivery, removal and Installation of fume cabinets complete with ancilliaries
2.1.1 | 1.5m free flow fume cupboard full polypropylene construction. Unit complete with fully bunded base constructed from laboratory grade phenolic resin G-Lab, black in colour. Unit with integrated lighting, vertical sliding polycarbonate 6mm sash on counter weight. Baffle to be removable for easy cleaning. External size 1500 x 750 x 1520mm excluding stand, overall height with stand@ 2420mm Internal working space 1200 x 630mm. Unit will have removable side service panels as well as top panel. | No | 3
2.1.2 | Fume cupboard light switch | No | 3 | Incl. | 0
2.1.3 | 15A plugs | No | 6 | Incl. | 0
2.1.4 | Broen water front control valve to operate with water outlet nozzle inside the unit. Complete with polypropylene drip cup and fittings | No | 3 | Incl. | 0
2.1.5 | Fume cupboard stand of 32 x 32mm epoxy coated tubular frames complete with connecting tie bars coated to client’s choice | No. | 3 | Incl. | 0
2.1.6 | Set under bench cabinet under the unit consisting of one 900mm wide and one 500mm wide units | No | 3 | Incl. | 0
2.1.7 | Polypropylene fume extraction ducting manifolds to connect 1.5m fume cupboards to chemical extractor fan on the roof including all bends, segments, bracketing and flashing |  |  | Sum | 27500
2.1.9 | CF 350/14/380V PVC centrifugal extractor fan complete with ducting all in rigid Polyprop with 1.5kw 380v motor with weather proof isolator | No | 3
2.1.10
2.1.11 | Supply, deliver and install
2.1.12 | Supply, deliver and install of plantroom equipment, vacuum pumps, LP Gas Manifold and Compressors
2.1.13 | Oil free 2 X Vacuum Pump, one duty, one standby, Flow Rate: 12m3/hr, Power Consumption: 0.45kW, 1PH, 240V, 50Hz for the supply of the vacuum air at each fume cabinet, inclusive of the copper reticulation | No. | 2
2.1.14 | Vacuum Receiver Tank, Capacity: 500L | No. | 1 | Incl.
2.1.15 | Vacuum Control Panel | Sum | 1 | Incl.
2.1.16 | Pressure Gauge | No. | 1 | Incl.
2.1.17 | Vacuum Trap | No. | 1 | Incl.
2.1.18 | Isolation Valves | No. | 4 | Incl.
2.1.19 | 2 X Oil Free Piston type compressor, one duty, one standby, Flowrate: 5m3/hr Power Consumption: 0.3kW, 1PH, 240V, 50Hz" for the supply of the compressed air at each fume cabinet, inclusive of the copper reticulation | No. | 2
2.1.20 | Refrigerant Type Compressor Air Drier, one duty, one standby, Refrigerant Type: R410A, Flowrate: 2.25L/s ,Pressure: 700kPa, Power Consumption: 1.62KW, 1PH, 240V, 50Hz | No. | 2 | Incl.
2.1.21 | Compressor Air Receiver Tank, Capacity: 1000L | No. | 1 | Incl.
2.1.22 | Oil Trap Filter | No. | 1 | Incl.
2.1.23 | Water Trap Filter | No. | 1 | Incl.
2.1.24 | Isolation Valves | No. | 4 | Incl.
2.1.25
2.1.26 | 2 x 8 LPG Manifold | No. | 2
2.1.27
2.1.29 | Supply, delivery and installation of high purity equipement, complete with all interconnecting 316L Stainless Steel Piping and fittings
2.1.30
2.1.31 | High purity installation for Argon/C2H2/N2O/Synthetic Air supply | No | 1
2.1.32
2.1.33 | LP Gas, vacuum and compressed air pipes as per SABS 460 class 2 pipe line
2.1.34
2.1.35 | 35mm copper pipe for reticulation network | m | 62
2.1.36
2.1.37 | 35mm 90 deg elbow | No. | 5
2.1.38
2.1.39 | 35mm x 35mm x 22mm Tee | No. | 4
2.1.40
2.1.41 | 35mm x 35mm x 15mm Tee | No. | 1
2.1.42
2.1.43 | 35mm x 35mm x 10mm Tee | No. | 1
2.1.44
2.1.45 | 22mm copper pipe for reticulation network | m | 24
2.1.46
2.1.47 | 22mm 90 deg elbow | No. | 9
2.1.48
2.1.49 | 22mm x 22mm x 15mm Tee | No. | 8
2.1.50
2.1.51 | 22mm x 22mm x 10mm Tee | No. | 3
2.1.52
2.1.53 | 15mm copper pipe for reticulation network | m | 32
2.1.54
2.1.55 | 15mm 90 deg elbow | No. | 12
2.1.56
2.1.57 | 15mm x 15mm x 15mm Tee | No. | 10
2.1.58
2.1.59 | 15mm x 15mm x 10mm Tee | No. | 4
TOTAL
TOTAL
2.1.60 | 10mm x 10mm x 10mm Tee | No. | 1
2.1.61 | 35mm lockable isolating valves | No. | 1
2.1.62 | 22mm Lockable Isolating Valves | No. | 1
2.1.63 | 15mm Lockable Isolating Valves | No. | 1
2.1.64 | 10mm Lockable Isolating Valves | No. | 5
2.1.65 | 48 kg LP Gas cylinders | No | 5
2.1.66 | LP Gas Manifold | No | 1
2.1.67 | 25mm LP Gas Regulators | No | 1
2.1.68 | 10mm aluminium flexible hoses | No | 8
2.1.69 | Water pipes as per SABS 460 class 1 pipe line
2.1.70 | 35mm copper pipe for reticulation network | m | 31
2.1.71 | 35mm 90 deg elbow | No. | 11
2.1.72 | 35mm x 35mm x 22mm Tee | No. | 10
2.1.73 | 35mm x 35mm x 15mm Tee | No. | 4
2.1.74 | 35mm x 35mm x 10mm Tee | No. | 3
2.1.75 | 22mm copper pipe for reticulation network | m | 12
2.1.76 | 22mm 90 deg elbow | No. | 5
2.1.77 | 22mm x 22mm x 15mm Tee | No. | 4
2.1.78 | 22mm x 22mm x 10mm Tee | No. | 2
2.1.79 | 15mm copper pipe for reticulation network | m | 16
2.1.80 | 15mm 90 deg elbow | No. | 6
2.1.81 | 15mm x 15mm x 15mm Tee | No. | 5
2.1.82 | 15mm x 15mm x 10mm Tee | No. | 2
2.1.83 | 10mm copper pipe for reticulation network | m | 30
2.1.84 | 10mm 90 deg elbow | No. | 11
2.1.85 | 10mm x 10mm x 10mm Tee | No. | 9
2.1.86 | 35mm lockable isolating valves | No. | 1
2.1.87 | 22mm Lockable Isolating Valves | No. | 1
2.1.88 | 15mm Lockable Isolating Valves | No. | 1
2.1.89 | 10mm Lockable Isolating Valves | No. | 3
2.1.90 | Drainage pipes as per Class 6 PVC pipe line
2.1.91 | Supply, deliver, off-load, store on site and install of Class 6 PVC piping including all fittings and other accessories for the drainage plumbing work , as required
2.1.92 | 110mm diameter | No | 25
2.1.93 | 50mmdiameter | No | 28
2.1.94 | 35 diameter | No | 7
2.1.95 | 28 diameter | No | 6
2.1.96 | 22 diameter | No | 18
2.1.97 | Supply, delivery, installation and commissioning of 50mm bend wiith Cleaning eye heel or (Rodding eye) | No | 2
2.1.98 | Supply, delivery, installation and commissioning of gully trap complete excluding building work (Gully Base by Main contractor) | No | 3
2.1.99 | Supply, delivery, installation and commissioning of 110mm Vent Valve | No | 3
2.1.100 | BENCHES AND GRANITE TABLES
2.1.101 | Replacement of the shelving and work benches (steel structure and granite tops) and shelving (stainless steel) in Mining and Environmental Geology Labs. Replacement of the shelving and work benches (steel structure and granite tops) and shelving (stainless steel) in the Hydrology and Water Resources Lab. | Sum |  | 45000 | 45000
2.1.102 | 180 degree double LPG bench taps. 180 degree double vacuum bench taps. All upright support from ceiling height to island benches.All interconnecting copper piping. e. All testing and commissioning. | Sum |  | 42500 | 42500
2.1.103 | ICP-OES Equipment - SUPPLY 1X2 MANIFOLD FOR ARGON WITH DOUBLE-STAGE REGULATOR, PIGTAILS, BULLNOSE AND NUTS - SUPPLY 40M OF ¼”S/STEEL PIPING - SUPPLY FITTINGS - SUPPLY 1 POINT INSIDE THE LAB - SUPPLY SIGNAGE, PIPELINE DECALS - PRESSURE TEST, COC | Sum |  | 75000 | 75000
2.1.104 | AAS Equipment Equipment - SUPPLY 1X1 ACETYLENE DOUBLE MANIFOLD COMPLETE WITH REGULATOR, FLASHBACK, PIGTAILS, BULLNOSE AND NUTS - SUPPLY 1X1 NITROUS OXIDE MANIFOLD COMPLETE WITH DOUBLE STAGE REGULATOR, PIGTAILS, BULLNOSE AND NUTS - SUPPLY 1X1 COMPRESSED AIR MANIFOLD COMPLETE WITH DOUBLE STAGE REGULATOR, PIGTAILS, BULLNOSE AND NUTS - SUPPLY 120M OF ¼” STAINLESS STEEL PIPING - SUPPLY 3 POINTS INSIDE THE LAB - SUPPLY FITTINGS - SUPPLY SIGNAGE.PIPELINE DECALS - PRESSURE TEST, COC - EARTHING OF ACETYLENE MANIFOLD SANS0108 | Sum |  | 72500 | 72500
Carried to summary

## Agricultural Sciences
FUME CABINETS, AGRICULTURAL SCIENCES DEPARTMENT, UNIVERSITY OF VENDA - MECHANICAL WORKS
ITEM | DESCRIPTION | UNIT | QTY | RATE | AMOUNT
3 | AGRICULTURAL SCIENCES FUME CABINETS
3.1 | Supply, delivery, removal and Installation of fume cabinets complete with ancilliaries
3,1,1 | 1.5m free flow fume cupboard full polypropylene construction. Unit complete with fully bunded base constructed from laboratory grade phenolic resin G-Lab, black in colour. Unit with integrated lighting, vertical sliding polycarbonate 6mm sash on counter weight. Baffle to be removable for easy cleaning. External size 1500 x 750 x 1520mm excluding stand, overall height with stand@ 2420mm Internal working space 1200 x 630mm. Unit will have removable side service panels as well as top panel. | No | 4
3,1,2 | Fume cupboard light switch | No | 4 | Incl.
3,1,3 | 15A plugs | No | 8 | Incl.
3,1,4 | Broen water front control valve to operate with water outlet nozzle inside the unit. Complete with polypropylene drip cup and fittings | No | 4 | Incl.
3,1,5 | Fume cupboard stand of 32 x 32mm epoxy coated tubular frames complete with connecting tie bars coated to client’s choice | No. | 4 | Incl.
3,1,6 | Set under bench cabinet under the unit consisting of one 900mm wide and one 500mm wide units | No | 4 | Incl.
3,1,7 | Polypropylene fume extraction ducting manifolds to connect 1.5m fume cupboards to chemical extractor fan on the roof including all bends, segments, bracketing and flashing |  |  | Sum | 27500
3,1,8 | CF 250/14/380V PVC centrifugal extractor fan complete with ducting all in rigid Polyprop with 0.55kw 380v motor with weather proof isolator | No | 3
3,1,9 | CF 350/14/380V PVC centrifugal extractor fan complete with ducting all in rigid Polyprop with 1.5kw 380v motor with weather proof isolator | No | 1
3,1,10 | Supply, deliver and install
3,1,11 | Supply, deliver and install of plantroom equipment, vacuum pumps, LP Gas Manifold and Compressors
3,1,12 | Oil free 2 X Vacuum Pump, one duty, one standby, Flow Rate: 12m3/hr, Power Consumption: 0.45kW, 1PH, 240V, 50Hz for the supply of the vacuum air at each fume cabinet, inclusive of the copper reticulation | No. | 2
3,1,13 | Vacuum Receiver Tank, Capacity: 500L | No. | 1 | Incl.
3,1,14 | Vacuum Control Panel | Sum | 1 | Incl.
3,1,15 | Pressure Gauge | No. | 1 | Incl.
3,1,16 | Vacuum Trap | No. | 1 | Incl.
3,1,17 | Isolation Valves | No. | 4 | Incl.
TOTAL
TOTAL
3,1,18 | 2 X Oil Free Piston type compressor, one duty, one standby, Flowrate: 5m3/hr Power Consumption: 0.3kW, 1PH, 240V, 50Hz" for the supply of the compressed air at each fume cabinet, inclusive of the copper reticulation | No. | 2
3,1,19 | Refrigerant Type Compressor Air Drier, one duty, one standby, Refrigerant Type: R410A, Flowrate: 2.25L/s ,Pressure: 700kPa, Power Consumption: 1.62KW, 1PH, 240V, 50Hz | No. | 2 | Incl.
3,1,20 | Compressor Air Receiver Tank, Capacity: 1000L | No. | 1 | Incl.
3,1,21 | Oil Trap Filter | No. | 1 | Incl.
3,1,22 | Water Trap Filter | No. | 1 | Incl.
3,1,23 | Isolation Valves | No. | 4 | Incl.
3,1,24
3,1,25 | Compressed air manifold | No. | 2
3,1,26
3,1,27 | 2 x 8 LPG Manifold | No. | 2
3,1,28
3,1,29 | Fume Cupboard Connections for LPG, Vacuum and air | Sum
3,1,30
3,1,31 | Receiver mount vacuum pump unit | No. | 1
3,1,32
3,1,33 | 2 x 3 compressed air manifold | No. | 2
3,1,34
3,1,35 | Fume Cupboard Connections for LPG, Vacuum and air | Sum | 4
3,1,36
3,1,37 | Receiver mount vacuum pump unit | No. | 1
3,1,38
3,1,39 | Compressed Air Manifold | No. | 1
3,1,40
3,1,41 | LPG Manifold | No. | 1
3,1,42
3,1,43 | Fume Cupboard Connections for LPG, Vacuum and air | Sum
3,1,44
3,1,45 | Supply, delivery and installation of high purity equipement, complete with all interconnecting 316L Stainless Steel Piping and fittings
3,1,46
3,1,47 | High purity installation for Argon/C2H2/N2O/Synthetic Air supply | No
3,1,48
3,1,49 | Testing and commissioning | Sum
3,1,50
3,1,51 | LP Gas pipes as per SABS 460 class 2 pipe line
3,1,52
3,1,53 | 35mm copper pipe for reticulation network | m | 124
3,1,54
3,1,55 | 35mm 90 deg elbow | No. | 44
3,1,56
3,1,57 | 35mm x 35mm x 22mm Tee | No. | 38
TOTAL
TOTAL
2.1.9 | 35mm x 35mm x 15mm Tee | No. | 13
2.1.10 | 35mm x 35mm x 10mm Tee | No. | 12
2.1.11 | 22mm copper pipe for reticulation network | m | 48
2.1.12 | 22mm 90 deg elbow | No. | 4
2.1.13 | 22mm x 22mm x 15mm Tee | No. | 3
2.1.14 | 22mm x 22mm x 10mm Tee | No. | 1
2.1.15 | 15mm copper pipe for reticulation network | m | 64
2.1.16 | 15mm 90 deg elbow | No. | 23
2.1.17 | 15mm x 15mm x 15mm Tee | No. | 20
2.1.18 | 15mm x 15mm x 10mm Tee | No. | 7
2.1.19 | 10mm copper pipe for reticulation network | m | 118
2.1.20 | 10mm 90 deg elbow | No. | 42
2.1.20 | 10mm x 10mm x 10mm Tee | No. | 35
2.1.21 | 35mm lockable isolating valves | No. | 2
2.1.22 | 22mm Lockable Isolating Valves | No. | 2
2.1.22 | 15mm Lockable Isolating Valves | No. | 2
2.1.23 | 10mm Lockable Isolating Valves | No. | 2
2.1.24 | 48 kg LP Gas cylinders | No | 10
2.1.25 | LP Gas Manifold | No | 1
2.1.26 | 25mm LP Gas Regulators | No | 2
2.1.27 | 10mm aluminium flexible hoses | No | 16
Water pipes as per SABS 460 class 1 pipe line
2.2.6 | 35mm copper pipe for reticulation network | m | 62
2.2.7 | 35mm 90 deg elbow | No. | 22
2.2.8 | 35mm x 35mm x 22mm Tee | No. | 19
2.2.9 | 35mm x 35mm x 15mm Tee | No. | 7
2.2.10 | 35mm x 35mm x 10mm Tee | No. | 6
2.2.11 | 22mm copper pipe for reticulation network | m | 24
2.2.12 | 22mm 90 deg elbow | No. | 9
2.2.13 | 22mm x 22mm x 15mm Tee | No. | 8
2.2.14 | 22mm x 22mm x 10mm Tee | No. | 3
2.2.15 | 15mm copper pipe for reticulation network | m | 32
2.2.16 | 15mm 90 deg elbow | No. | 12
2.2.17 | 15mm x 15mm x 15mm Tee | No. | 10
2.2.18 | 15mm x 15mm x 10mm Tee | No. | 4
2.2.19 | 10mm copper pipe for reticulation network | m | 59
2.2.20 | 10mm 90 deg elbow | No. | 21
TOTAL
TOTAL
2.2.21 | 10mm x 10mm x 10mm Tee | No. | 18
2.2.22 | 35mm lockable isolating valves | No. | 1
2.2.23 | 22mm Lockable Isolating Valves | No. | 1
2.2.24 | 15mm Lockable Isolating Valves | No. | 1
2.2.25 | 10mm Lockable Isolating Valves | No. | 5
2.2.26 | Drainage pipes as per Class 6 PVC pipe line
2.2.27 | Supply, deliver, off-load, store on site and install of Class 6 PVC piping including all fittings and other accessories for the drainage plumbing work , as required
2.2.28 | 110mm diameter | No | 50
2.2.29 | 50mmdiameter | No | 55
2.2.30 | 35 diameter | No | 13
2.2.31 | 28 diameter | No | 12
2.2.32 | 22 diameter | No | 35
2.2.33 | Supply, delivery, installation and commissioning of 50mm bend wiith Cleaning eye heel or (Rodding eye) | No | 4
2.2.34 | Supply, delivery, installation and commissioning of gully trap complete excluding building work (Gully Base by Main contractor) | No | 5
2.2.35 | Supply, delivery, installation and commissioning of 110mm Vent Valve | No | 6
2.2.36 | New stainless-steel shelving at the analytical lab including ventilation. Replacement of the shelving and work benches (steel structure and granite tops) and shelving (stainless steel) in the Animal Science Lab Removal the work benches in one common lab to create better circulation | Sum |  | 45000 | 45000
2.2.37 | 180 degree double LPG bench taps. 180 degree double vacuum bench taps. All upright support from ceiling height to island benches.All interconnecting copper piping. e. All testing and commissioning. | Sum |  | 42500 | 42500
Carried to summary